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PART 1 GENERAL

1.1 Section includes:
A. Sensor that continuously measures conductivity and/or resistivity in aqueous solutions.
1.2 Measurement Procedures
A. The method of measurement will be contacting conductivity and/or resistivity with a probe or sensor
1.3 Alternates

A. Probes or sensors that do not communicate with Hach model sc200 or Polymetron 9500 controllers will not be acceptable.

1.4 System Description

A. Performance Requirements for 8398.2 Sensor
1. Cell Constant: 2.35 cm-1
2. Measurement Range: 0 to 2000 mS/cm 

B. Performance Requirements for 8398.3 Sensor
1. Cell Constant: 2.35 cm-1
2. Measurement Range: 0 to 2000 mS/cm 


C. Performance Requirements for 8398.5 Sensor
1. Cell Constant: 2.35 cm-1
2. Measurement Range: 0 to 2000 mS/cm 
 

1.5 Certifications
A. EMC: CE compliant for conducted and radiated emissions CISPR 11 (Class A limits), EMC Immunity EN 61326-1 (Industrial limits) when part of an approved system
B. Safety: General Purpose UL/CSA 61010-1 with cETLus safety mark when part of an approved system
C. Australian C-TICK and Korean KC Markings when part of an approved system
1.6 Environmental Requirements

A. Operational Criteria

1. Maximum Operating Temperature: 60 °C at 2 bar (140 °F at 29 psi)
2. Maximum Sample Temperature: 140 °C at 18 bar (284 °F at 261 psi)

3. Maximum Pressure: 18 bar at 140 °C (261 psi at 284 °F)
4. Flow Rate: 20L/hr minimum (60 L/hr ideal)

5. Relative humidity: 10 to 95%, non-condensing
1.7 Maintenance Service

A. Scheduled Maintenance
1. Annually

a. Temperature

1) Calibration

b. Measurement

1) Calibration

B. Unscheduled Maintenance
1. Clean sensor with appropriate cleaning solution at appropriate intervals depending on the sample tested. 
PART 2 PRODUCTS

2.1 Manufacturer
A. Hach Lange Sárl, Geneva, Switzerland
2.2 Manufactured Unit

A. The 8398 series conductivity sensor consists of:

1. A contacting conductivity sensor composed of PEEK.
2. Integral Cable

3. The sensor is individually factory-tested to determine its absolute cell constant and temperature constant for accuracy.
2.3 Equipment
A. The Polymetron 8398 Conductivity Sensors work with Hach model sc200 or Polymetron 9500 controllers only.
B. The sensor has a built in Pt 100 temperature sensor.
C. Wetted materials as follows:

1. 8398.2 Sensor
a. PEEK
2. 8398.3 Sensor

a. PEEK
3. 8398.5 Sensor

a. PEEK
b. EPDM or Viton Gaskets
2.4 Components

A. Standard equipment
1. Sensor

2. Integral Cable

B. Dimensions: 
1. 8398.2 Sensor: 40 mm x 83.5 mm (1.57 in x 3.29 in)

2. 8398.3 Sensor: 40 mm x 88.5 mm (1.57 in x 3.49 in)

3. 8398.5 sensor: 40 mm x 95.5 mm (1.57 in x 3.76 in)
C. Weight: 
1. 8398.2 Sensor: 0.7 kg (1.6 lbs)
2. 8398.3 Sensor: 0.9 kg (2.0 lbs)
3. 8398.5 Sensor: 0.9 kg (2.0 lbs)
2.5 Optional Accessories
A. Cables

1. 5m (16 ft)

2. 10 m (33 ft)

3. 15 m (50 ft)

4. 20 m (66 ft)

B. Application specific mounting hardware

PART 3 EXECUTION

3.1 Preparation
A. The sensor must be mounted to a Hach or Polymetron mounting assembly directly in the solution to be measured.
3.2 Installation

A. Contractor will install the analyzer in strict accordance with the manufacturer’s instructions and recommendation.

B. Manufacturer’s representative will include a half-day of start-up service by a factory-trained technician, if requested.

1. Contractor will schedule a date and time for start-up.

2. Contractor will require the following people to be present during the start-up procedure. 

a. General contractor

b. Electrical contractor

c. Hach Company factory trained representative

d. Owner’s personnel

e. Engineer
3.3 Manufacturer’s Service and Start-Up

A. Contractor will include the manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

B. Contractor will include a manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment startup through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

C. Items A and B are to be performed by manufacturer’s factory-trained service personnel.  Field service and factory repair by personnel not employed by the manufacturer is not allowed.

D. Use of manufacturer’s service parts and reagents is required.  Third-party parts and reagents are not approved for use. 
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