PART 1  GENERAL

  - Section includes

    - Amperometric ozone analyzer for the measurement of ozone in water.

  - Measurement Procedures

    - The method of measuring ozone will be with two-electrode amperometric sensor with membrane selective to ozone, immersed into an electrolytic medium.

  - Alternates

    - Other methods of ozone measurement, such as single electrode amperometric, which can show interference from chlorine, are not acceptable.

    - Direct amperometric method of measurement is not acceptable.

    - Systems that do not use a gas diffusion membrane are not acceptable.

    - Systems that are affected by low conductivity levels are not acceptable.

    - Systems that are not capable of measuring ozone in pure water applications are not acceptable.

  - System Description

    - Performance Requirements

      - Measurement range: 0 to 2 ppm ozone

      - Detection limit: 5 ppb or 0.005 mg/L ozone

      - Accuracy: 3% or ±10 ppb ozone, whichever is greater

      - Response time: 90% in less than 90 seconds

      - Calibration

        - Electrical zero or chemical zero with dechlorinated or deozoned water

        - Calibration of the slope by comparison with laboratory instrument

        - pH calibration with single or two point using standards or comparison with lab method and process sample

  - Certifications (when connected to an sc controller)

    - EMC: CE compliant for conducted and radiated emissions CISPR 11 (Class A limits), EMC Immunity EN 61326-1 (Industrial limits)

    - Safety: General Purpose UL/CSA 61010-1 with cETLus safety mark

    - Australian C-TICK and Korean KC Markings

    - NEMA 4X/IP66 dust and water ingress ratings

  - Environmental Requirements

    - Operational Criteria

      - Storage Temperature Range

        - -20 to 60 °C (-4 to 140 °F)

      - Operating Temperature Range

        - 0 to 45 °C (32 to 113 °F)

      - Sample Temperature Range

        - 2 to 45°C (35 to 113 °F)

      - Sample Pressure Range

        - 1.4 to 28 psi (0.1 to 2 bar)

      - Sample pH Range

        - pH 4 to pH 8 (Acidification unit available for >8 pH)

  - Warranty

    - The product includes a one-year warranty from the date of shipment.

  - Maintenance Service

    - Scheduled Maintenance

      - Ozone sensor calibration every two months

      - Replace tubing annually

    - Unscheduled Maintenance

      - Replace membrane and electrolyte each 6 months (3 to 12 month range depending on sample)

      - Clean electrode and pH sensor flow cell as required by sample (recommended every 6 months)

PART 2  PRODUCTS

  - Manufacturer

    - Hach Company, Loveland, CO

      - Model 9185sc Amperometric Ozone Analyzer

  - Manufactured Unit

    - The 9185sc ozone analyzer consists of:

      - Two-electrode amperometric ozone sensor

      - Ozone sensor flow cell

      - Mounting panel

      - Digital cable to connect analyzer to sc controller

  - Equipment

    - The 9185sc ozone analyzer works with Hach sc controllers only and connects to the controller by a digital plug-and-play interface.

    - The amperometric cell of the analyzer consists of:

      - Gold cathode

      - Silver anode

      - pH buffered Potassium Chloride reference electrolyte

      - Sensor membrane to filter chlorine species selectively and to provide interface between the electrochemical cell and the sample

    - The 9185sc sensor automatically compensates for temperature via integral temperature sensor.

    - Wetted materials as follows:

      - Ozone electrode: gold cathode/silver anode

      - Sensor body: PVC

      - Measuring cell: Acrylic

  - Components

    - Standard Equipment

      - Mounting Panel

      - Amperometric Ozone sensor with membrane and electrolyte

      - Flow cell

      - Digital sc sensor cable

      - User Manual

    - Dimensions: 270 x 250 mm x 155 mm (10.63 x 9.84 x 6.1 in)

    - Weight: 6.5 pounds (14.3 kg)

  - Accessories

    - Required

      - Hach sc controller

    - Optional

      - Acidification Unit

      - Intermittent Flow Unit

PART 3  EXECUTION

  - Preparation

      - Mounting

        - Instrument is able to be wall or panel mounted. The pre-assembled analyzer panel must be mounted to allow clearance above the measurement sensor for sensor removal.

      - Sample inlet

        - ¼ inch OD connection

      - Sample Flow Rate

        - Minimum of 14L/hr

      - Drain Fitting

        - ½ inch ID connection

  - Installation

    - Contractor will install the analyzer in strict accordance with the manufacturer’s instructions and recommendation.

    - Manufacturer’s representative will include a half-day of start-up service by a factory-trained technician, if requested.

      - Contractor will schedule a date and time for start-up.

      - Contractor will require the following people to be present during the start-up procedure.

        - General contractor

        - Electrical contractor

        - Hach Company factory trained representative

        - Owner’s personnel

        - Engineer

  - Manufacturer’s Service and Start-Up

    - Contractor will include the manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

    - Contractor will include a manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment startup through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

    - Items A and B are to be performed by manufacturer’s factory-trained service personnel.  Field service and factory repair by personnel not employed by the manufacturer is not allowed.

    - Use of manufacturer’s service parts and reagents is required.  Third-party parts and reagents are not approved for use.

- END OF SECTION