



Automatic Portable Sampler
An automatic sampler to collect and store unbiased representative water samples for laboratory analysis. A portable sampler’s lightweight design, low power requirements, and mobility make it ideal for field applications. Equipped with  a peristaltic pump with spring-mounted rollers and a standard liquid detect sensor for sample volume accuracy. A USB
connection and an auxiliary port for pulse or 4-20 inputs are standard. Factory-installed options include sensor ports for integrated water quality, level or flow, a port for Modbus RS485 communication or a rain sensor and an alternate non- contact liquid detect sensor. Three base sizes to accommodate different requirements; compact for smaller manhole applications, standard for a majority of applications and large to allow for large bottle composite sampling. Using appropriate modules, inputs and outputs can be extended for a single high-voltage relay or a comprehensive set of analogue outputs or analogue and digital outputs and relays.
Technical Data
	  Sampling Features

	Programs
	Dual Programs: Up to 2 sample programs can be run sequentially, in parallel, or according to day of week scheduling; enabling a single sampler to function like multiple samplers

	Sampling mode
	Pacing: Time weighted, Flow weighted, Time table, Flow table, Event
Distribution: Single bottle composite, multi-bottle composite, multi-bottle discrete, bottles per sample, samples per bottle or a combination of bottles per sample and samples per bottle

	Operating mode
	Continuous or non-continuous

	Status messages
	Communicates what program is running, if there are any missed samples, when the next sample will be taken, how many samples remain, number of logged channels, time of last measurement, memory available, number of active channels, if alarms were triggered, when alarms were triggered, active sensors and cabinet temperature

	Alarm
	Configurable alarms that show on status screen and are recorded in diagnostics alarm logs. Alarms can be set for system diagnostics and logging such as program end, sample complete, missed samples and full bottle. Channel alarms are setpoint alarms for the recorded measurements (channels), such as pH, level and power supply voltage.

	Manual sampling
	Initiates a sample collection independent of program in progress

	Automatic shutdown
	Multiple Bottle Mode: After complete revolution of distributor arm (unless Continuous Mode is selected)
Composite Mode: After preset number of samples have been delivered to composite container, from 1 to 999 samples, or upon full container.

	Sample volume
	Programmable in 10 mL increments from 10 to 10,000 mL

	Sample interval
	Selectable in single increments from 1 to 9999 flow pulses or 1 to 9999 mins in one minute increments

	Sample trigger
	When equipped with flow sensor or pH/temperature sensor or peripheral monitoring options, sampling can be triggered upon an upset condition when field selectable limits are exceeded

	Data logging
	SAMPLE HISTORY - Stores up to 4000 entries for sample time stamp, bottle number and sample status (success, bottle full, rinse error, user abort, distributor error, pump fault, purge fail, sample timeout, power fail and low main battery).
MEASUREMENTS - Stores up to 325,000 entries for selected measurement channels in accordance with the selected logging interval.
EVENTS - Ability to store up to 2000 entries in Sample History logging. Records Power On, Power Fail, Firmware Updated, Pump Fault, Distributor Arm Error, Low Memory Battery, Low Main Battery, User On, User Off, Program Started, Program Resumed, Program Halted, Program Completed, Grab Sample, Tube Change Required, Sensor Communication Errors, Cooling Failed, Heating Failed, Thermal Error Corrected.

	Diagnostics
	View event and alarm logs as well as maintenance diagnostics


	   Portable Sampler

	Housing material
	Impact resistant ABS plastic, 3-section construction Double-walled base with 2.54 cm insulation, direct ice contact with bottles (Sampler)

	Sample container
	STANDARD BASE CAPACITY:
(24) 1 L Polyethylene or 350 mL Glass bottles
(8) 2.3 L Polyethylene or 1.9 L Glass bottles
(4) 3.8 L Polyethylene or 3.8 L Glass bottles
(2) 3.8 L Polyethylene or 3.8 L Glass bottles
(1) 21 L Polyethylene composite bottle or
(1) 15 L Polyethylene composite bottle or
(1) 20 L Polyethylene composite bottle or
(1) 10 L Polyethylene or 10 L Glass composite bottle
COMPACT BASE CAPACITY:
(24) 575 mL Polyethylene bottles
(8) 950 mL Glass bottles
(1) 10 L Polyethylene or 10 L Glass bottle
COMPOSITE BASE CAPACITY:
(1) 21 L Polyethylene bottle
ICE CAPACITY:
Compact Base: 3.9 kg with (24) 575 mL PE bottles Standard Base: 14.5 kg with (24) 350 mL Glass bottles

	Temperature
	Operating: 0 - 49 °C
Storage: -40 - 60 °C

	Power requirements (Voltage)
	12 V DC

	Dimensions
	Standard Base: 50.5 cm x 69.4 cm Compact Base: 44.1 cm x 61 cm Composite Base: 50.28 cm x 79.75 cm

	Weight
	Standard:
15 kg with (24) 1 L Polyethylene bottles; 14.8 kg with 10 L Polyethylene bottle
Compact Base:
12.2 kg with (24) 575 mL Polyethylene bottles; 12.9 kg with 10 L Polyethylene bottle
Composite Base:
15 kg with (1) 21 L Polyethylene bottle

	   Controller

	Housing material
	PC/ABS blend, NEMA 4X, 6, IP68, corrosion and ice resistant (Controller)

	Display
	1/4 VGA, colour; self-prompting/menu-driven program

	User interface
	Membrane switch keypad with 2 multiple function soft keys

	Languages user interface
	en, fr, es, it, de, pt, zh, tr, hu, cz, pl, ro, hr, el, sl, sk, fi, ru, ja, ko

	Lock function
	Access code protection prevents tampering

	Memory
	Sample history: 4000 records; Data log: 325,000 records; Event log: 2000 records

	Communication capabilities
	USB and optional RS485 (Modbus)

	Inputs
	One 0/4-20 mA input for flow pacing

	Certifications
	CE, UL


	   Sample Pump and Strainer

	Sample pump
	High-speed peristaltic, dual roller, with 0.95 ID x 1.6 OD cm (3/8” ID x 5/8” OD) pump tube

	Vertical lift
	8.5 m using 8.8 m maximum of 3/8” Vinyl intake tube at sea level at 20 to 25 °C

	Tubing
	Pump tubing: 9.5 mm ID x 15.9 mm OD Silicone
Intake tubing: 1.0 - 4.75 m minimum length, 1/4” or 3/8” ID Vinyl or 3/8” ID PTFE-lined Polyethylene with protective outer cover (black or clear)

	Sample volume repeatability
	Typical: ±5% of 200 mL sample volume with: 4.6 m vertical lift, 4.9 m of 3/8” Vinyl intake tube, single bottle, full bottle shut-off at room temperature and 1524 m elevation

	Sample volume accuracy
	Typical: ±5% of 200 mL sample volume with: 4.6 m vertical lift, 4.9 m of 3/8” Vinyl intake tube, single bottle, full bottle shut-off at room temperature and 1524 m elevation

	Sample transport velocity
	0.9 m/s with 4.6 m vertical lift, 4.8 m of 3/8” Vinyl intake tubing, 21 °C and 1524 m elevation

	Pump flow rate
	4.8 L/min at 1 m vertical lift with 3/8” intake tube typical

	Internal clock
	±1 second per day at 25 °C

	Intake
	Strainers: Choice of PTFE and 316 stainless steel construction, or all 316 stainless steel in standard size, high velocity, and low profile for shallow depth applications
Purge: Air purged automatically before and after each sample; duration automatically compensates for varying intake line lengths
Rinse: Intake line automatically rinsed with source liquid prior to each sample, from 1 to 3 rinses
Retries or Fault: Sample collection cycle automatically repeated from 1 to 3 times if sample not obtained on initial attempt


Scope of delivery
[ ] (delete what is not needed)
[ ] Basic Controller (includes 4-20 aux port and USB port) [ ] Controller with sensor ports
[ ] Controller with Rain/RS485 port
[ ] Controller with sensor ports and Rain/RS485 port
Portable Base [ ] Compact
[ ] Standard [ ] Large
Power
[ ] 12 V Lead acid battery [ ] Battery charger
[ ] 100-120 V power supply US plug [ ] 230 V power supply EU plug
Bottle Sets
[ ] Composite 21 L Polyethylene bottle kit for Large Base [ ] Composite 10 L Polyethylene bottle kit
[ ] Discrete (24) 1 L Polyethylene bottle kit for Standard Base [ ] Discrete (24) 350 mL Glass bottle kit for Standard Base
[ ] Discrete (24) 575 mL Polyethylene bottle kit for Compact Base [ ] Composite 15 L Polyethylene bottle kit for Standard Base
[ ] Composite 10 L Glass bottle kit
Intake Tubing
[ ]  Vinyl - 7.5 m, 3/8” ID x 5/8” OD tubing and PTFE/stainless steel strainer [ ]  PTFE-lined polyethylene 7.5 m, 3/8” ID x 5/8” OD with connector kit
Sensors
[ ] Non oil-filled 0-3 m range, 9 m cable with connector [ ] Oil-filled 0-3 m range, 9 m cable with connector
[ ] Digital pH sensor
[ ] Digital pH sensor for Europe only [ ] Downlooking ultrasonic sensor
[ ] Suspended downlooking ultrasonic sensor [ ] In-pipe ultrasonic sensor
Accessories
[ ] Cable assy, aux half, 7P, 1.2 m [ ] Cable assy, aux half, 7P, 7.6 m
[ ] Universal junction box for 4-20 mA [ ] Input/Output 1 relay
[ ] Input/Output digital outputs/relays, analogue inputs
Hardware
[ ] Mounting hardware kit
[ ] Sampler suspension harness (use with manhole suspension bracket) [ ]  Manhole support bracket/spanner 18-27”
[ ]  Manhole support bracket/spanner 28-48” [ ]  Manhole support bracket 18-26”
Services
[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.
[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.
Brand:
Hach
Product:
AS950 Automatic Portable Sampler
PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.
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