GENERAL

  - Section includes:

    - Chlorine analyzer for monitoring of free or total residual chlorine

  - Measurement Procedures

    - The method of measuring free or total chlorine will be colorimetric.  Instrument chemistry will employ N, N-diethyl-p-phenylenediamine (DPD) method.

  - Alternates

    - Other methods of chlorine measurement such as amperometric, potentiometric, and iodometric that employ electrodes or other electrochemical techniques are not acceptable.

  - System Description

    - Performance Requirements

      - Measurement range:

        - 0 to 5 mg/L (ppm) free or total residual chlorine

      - Accuracy

        - ± 5% of reading or ±0.03 mg/L (ppm), whichever is greater

      - Precision

        - 5% of reading or 0.01 mg/L (ppm), whichever is greater

      - Minimum detection limit

        - 0.03 mg/L (ppm)

      - Resolution

        - 0.01 mg/L (ppm)

      - Repeatability

        - 0.05 mg/L (ppm)

      - Cycle Time

        - 2.5 minutes

  - Certifications

    - CE compliant for conducted and radiated emissions CISPR 11 (Class A limits), EMC Immunity EN 61326-1 (Industrial limits), and EN 61010-1

    - General Purpose UL/CSA 61010-1 with cETLus safety mark

    - IP62 dust and water ingress protection rating

    - Australian CTICK and Korean KC Marking

  - Environmental Requirements

    - Operational Criteria

      - Sample flow rate

        - 200 to 500 mL/minute

      - Sample pressure (without conditioning kit)

        - 1 to 5 psi (0.07 to 0.34 bar)

      - Sample pressure (with conditioning kit)

        - 120 psi (8.27 bar)

      - Sample temperature

        - 41 to 104 °F (5 to 40 °C)

      - Operating temperature

        - 41 to 104 °F (5 to 40 °C)

      - Operating humidity

        - 90% at 40 °C maximum

  - Warranty

    - The product includes a one-year warranty from the date of shipment

  - Maintenance Service

    - Scheduled Maintenance

      - Monthly

        - Reagent replacement

      - Annually

        - Analyzer tubing replacement

    - Unscheduled Maintenance

      - Pump tubing replacement is operating temperature dependent

        - Operating temperature below 80 °F: six-month intervals

        - Operating temperature above 80 °F: three-month intervals

PRODUCTS

  - Manufacturer

    - Hach Company, Loveland, CO

      - Model Cl17 Chlorine Analyzer, Free Chlorine Residual

      - Model Cl17 Chlorine Analyzer, Total Chlorine Residual

  - Manufactured Unit

    - The Cl17 Chlorine analyzer consists of a sample and reagent valve and pump, measurement cell, controller, and is shipped with buffer and indicator solutions.

  - Equipment

    - The analyzer must be housed in a NEMA 12 enclosure that is IP62 rated with the gasketed door latched.

    - The analyzer shall be capable of measuring free or total residual chlorine by changing the tubing and indicator and buffer solutions.

    - A measurement shall be taken every 2.5 minutes and results displayed by a three digit LCD readout in the range of 0 to 5 mg/L.

    - The analyzer must operate using 115V or 230V selectable AC power.

    - The analyzer must perform a self-test and auto-blanking between analysis points to compensate for sample color, turbidity, and changes in light intensity due to voltage fluctuations or light source aging.

    - The analyzer shall operate with an LED light source at a peak wavelength of 510nm.

    - The analyzer must be able to operate unattended for 30 days between chemical reagent changes and measurement cell cleaning.

    - The analyzer has two feed control (relay) operation modes to operate chemical feed pumps. Available control options are:

      - On/off control where the concentration alarm outputs activate or deactivate a pump when chlorine levels fall below or exceed acceptable levels.

      - Proportional control where the 4-20mA output current is scaled to pace a feed pump proportional to output.

    - The analyzer has standard optically isolated analog outputs, selectable as 0/4 to 20mA, field programmable over any portion of the analyzer range

    - The analyzer has two standard SPDT relay alarms, with contacts rated for 5 amp resistive loads at 230V AC power. Alarm options include concentration set point, analyzer system warning, and analyzer system shut down.

  - Components

    - Standard Equipment

      - Cl17 Free or Total Chlorine analyzer

      - One-Month Supply of reagents

      - Installation kit

      - Maintenance kit

      - Sample conditioning kit

        - Pressure regulator, strainer, and shut off valve

      - Wall mount kit

      - User manual

    - Dimensions:  13.5 x 17.9 x 7 inches (343 x 455 x 178 mm)

    - Shipping weight: 16 lbs (7.3 kg)

  - Optional Accessories

    - Power Cord

    - Maintenance kit with preassembled tubing

    - Pocket Colorimeter II for free and total chlorine (high and low range combination)

EXECUTION

  - Preparation

      - Mounting

        - The Cl17 Free or Total Chlorine analyzer can be wall mounted only.

      - Required Clearances

        - Horizontal: 15.2 in (386 mm), 26 inches (686 mm) ideal

        - Vertical: 19 inches (483 mm)

        - Depth: 20 inches (508 mm)

      - Sample inlet

        - 0.25 inch OD polyethylene tubing

      - Sample outlet

        - 0.50 inch ID flexible tubing

      - Overflow drain

        - 0.50 inch ID flexible tubing

      - Air purge quick connect

        - 0.25 inch OD polyethylene tubing (optional)

  - Installation

    - Contractor will install the analyzer in strict accordance with the manufacturer’s instructions and recommendation.

    - Manufacturer’s representative will include a half-day of start-up service by a factory-trained technician, if requested.

      - Contractor will schedule a date and time for start-up.

      - Contractor will require the following people to be present during the start-up procedure.

        - General contractor

        - Electrical contractor

        - Hach Company factory trained representative

        - Owner’s personnel

        - Engineer

  - Manufacturer’s Service and Start-Up

    - Contractor will include the manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

    - Contractor will include a manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment startup through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

    - Items A and B are to be performed by manufacturer’s factory-trained service personnel.  Field service and factory repair by personnel not employed by the manufacturer is not allowed.

    - Use of manufacturer’s service parts and reagents is required.  Third-party parts and reagents are not approved for use.

- END OF SECTION