PART 1	GENERAL

  - Section includes

    - Sensor that continuously measures conductivity and/or resistivity in aqueous solutions.

  - Measurement Procedures

    - The method of measurement will be contacting conductivity and resistivity with a probe or sensor.

  - Alternates

    - Probes or sensors that do not communicate with the Hach model sc100 or sc1000 digital controller will not be accepted.

  - System Description

    - Performance Requirements

      - Measurement range: 0.057 to 200,000 µS/cm

      - Accuracy: ±0.01% of reading, all ranges

  - Certifications

    - General Purpose CSA/CSANRTL and FM (UL Pending) when part of an approved system.

    - Class 1, Div 2 Groups A thru D CSA/CSANTRL and FM (UL Pending) when part of an approved system.

  - Environmental Requirements

    - Operational Criteria

      - Sample flow rate: 0 to 3 meters (0 to 10 feet) per second, maximum, fully immersed

      - Operating temperature: dependant on probe selection

      - Temperature and pressure limits, depending on materials. (Other mounting hardware or piping material may change the listed rating.):

        - Compression fitting sensor:

          - With Kynar® (PVDF) compression fitting: 302°F at 25 psi (150°C at 1.7 bar) or 97° at 150 psi (36°C at 10.3 bar)

          - With 316 stainless steel compression fitting: 302°F at 200 psi (150°C at 13.7 bar)

          - With ball valve assembly: 257°F at 150 psi (125°C at 10.3 bar)

        - Non-metallic general purpose sensor: 302°F at 100 psi (150°C at 6.8 bar) or 68° at 200 psi (20°C at

13.7 bar)

        - Boiler/condensate sensor:

          - No hardware: 212°F at 300 psi (100°C at 20.7 bar)

          - With integral polypropylene j-box head: 198°F at 300 psi (92°C at 20.7 bar)

          - With integral aluminum or 316 stainless steel j-box head: 392°F at 300 psi (200°C at 20.7 bar)

        - Sanitary (CIP) flange sensor:

          - Sensor only: 302°F at 100 psi (150°C at 6.8 bar)

          - With GLI sanitary mount hardware: 257°F at 150 psi (125°C at 10.3 bar)

      - Insertion depth: 102 mm (4 inches)

  - Warranty

    - The sensor is warranted for 1 year from date of purchase against material and workmanship.

  - Maintenance Service

    - Scheduled maintenance:

      - Clean to maintain measurement accuracy. Schedule (days, weeks, etc.) is affected by the characteristics of the process solution and should be determined by operating experience.

PART 2	PRODUCTS

  - Manufacturer

    - Hach Company, Loveland, CO

      - Hach 3400-SC Contacting Conductivity Sensor

  - Manufactured Unit

    - The Hach 3400-SC Contacting Conductivity Sensor consists of:

      - Conductivity probe with titanium or stainless steel electrodes

      - Integral cable

      - The sensor is individually factory-tested to determine its absolute four-digit cell constant and temperature constant for accuracy.

      - Digital gateway device

  - Equipment

    - The Hach 3400-SC sensor works with Hach models sc100 or sc1000 controllers only.

    - The probe has a built in Pt 1000 RTD temperature compensator.

    - Wetted materials as follows:

      - Compression fitting sensor:

        - Electrodes: titanium or 316 stainless steel outer electrode for extended sensor body style used with ball valve assembly)

        - Insulator: PTFE Teflon®

        - O-ring seals: treated Viton®

      - Non-metallic general purpose conductivity sensor:

        - Electrodes: graphite

        - Insulator: Ryton®

        - O-ring seals: Viton®

      - Boiler/condensate sensor:

        - Electrodes: 316 stainless steel and titanium

        - Insulator: PEEK®

        - O-ring seals: fluoroelastomer

      - Sanitary (CIP) flange sensor:

        - Electrodes: 316 stainless steel

        - Insulator: PTFE Teflon®

        - O-ring seals: Viton®

  - Components

    - Standard equipment:

      - Probe

      - Integral cable

      - Digital gateway

      - Manual

    - Dimensions: dependant on mounting configuration

    - Weight: dependant on mounting configuration

  - Accessories

    - Plug-in extension cables to extend the distance between the sensor and cable up to 1000 meters (3240 ft.).

    - Junction box for extension cables (6-position terminal strip supplied in integrally mounted junction box of polypropylene, aluminum, or 316 stainless steel). Must be used for lengths greater than 100 meters.

    - Extension cables

    - Mounting hardware

    - Compression fittings

    - Low-volume flow chamber

PART 3	EXECUTION

  - Preparation

    - The sensor must be mounted to a Hach mounting assembly directly in the solution to be measured.

  - Installation

    - Contractor will install the analyzer in strict accordance with the manufacturer’s instructions and recommendation.

    - Manufacturer’s representative will include a half-day of start-up service by a factory-trained technician, if requested.

      - Contractor will schedule a date and time for start-up.

      - Contractor will require the following people to be present during the start-up procedure.

        - General contractor

        - Electrical contractor

        - Hach Company factory trained representative

        - Owner’s personnel

        - Engineer

  - Manufacturer’s Service and Start-Up

    - Contractor will include the manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

    - Contractor will include a manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment startup through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

    - Items A and B are to be performed by manufacturer’s factory-trained service personnel.  Field service and factory repair by personnel not employed by the manufacturer is not allowed.

    - Use of manufacturer’s service parts and reagents is required.  Third-party parts and reagents are not approved for use.

END OF SECTION