# EM950 Sensor Firmware Release Notes

| Product Name / Model: | EM950.XXXXX |
| --- | --- |
| Version: | Application v1.04, Boot v1.00 |
| Date: | September 21, 2023 |
| Project No.: | RES0800 |
| Part No. / Board No.: | Firmware: 9074979, Programmed Part: 9074900, Ckt Bd: 9065600 |
| Technical Lead: | Fabian Kratz (OTT) |
| Target Processor: (Hach p/n, Mfg p/n) | Mfg: STM32F103CBT6, Hach: 9064900 |

Document Overview:

This document lists differences between a product and its predecessor version.  Differences are categorized as defects which have been corrected, improvements, and new features.  The tables used to list each category of differences are all structured identically.  Item numbers are sequentially numbered differences.  Tracking numbers refer to issue tracking systems.  This product uses {later} to track version differences.  External communication indicates whether the difference should be reported to external parties outside of Hach.  This column should be Y/N for Yes/No.  When an external communication item is marked “Y”, the description should be written for laymen (i.e., avoid technical jargon).

# Version 1.04

The release date of this version is September 21, 2023

## Defects

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |
| 1 |  | N | Correct wrong condition check which avoids using of more than three calibration coefficients as well as version 1.03 was failing in production floor test so this version has fix for that bug. |

# Version 1.03

The release date of this version is March 15, 2023.

# Defects

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |
| 1 |  | N | Correct wrong condition check which avoids using of more than three calibration coefficients |

## Improvements

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |
| 1 |  | N |  |

## New Features

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |
| 1 |  |  |  |

# Version 1.02 = 0.16

The release date of this version is March 15, 2013.

## Defects

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |
| 1 |  | N | Corrections were made to the unit of measure library for values falling below enumerated 130. The enumerated list issue was resolved and updated with radians unit of measure. |

## Improvements

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |
|  |  | N | Changed CFS to GPM units conversion constant for improved accuracy, and for synchronization with other projects’ TUnit.cpp file. |
|  |  | Y | Originally, the depth returned by the sensor was with respect to the bottom of the sensor enclosure. Jonathon Cronin and Christel Valentine requested the depth reference be at the centerline of the velocity electrode, so a mathematical modification was made in the pressure transducer driver to adjust the depth measurement by subtracting the mechanical offset between the face of the transducer and the centerline of the velocity electrode. |

## New Features

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |
| 1 |  |  |  |

# Version 1.01 = 0.15

The release date of this version is May 24, 2012.

## Defects

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |

## Improvements

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |
|  |  | N | Added velocity draw-down level correction. |
|  |  | N | Added mechanical offset from pressure transducer to bottom of sensor to correct level. |

## New Features

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |
| 1 |  |  |  |

# Version 0.14

The release date of this version is May 4, 2011.

## Defects

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |

## Improvements

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |
|  |  | N | Replaced the default linear level calibration with a polynomial for both gain and offset as a function of temperature. Note: this required the calibration file format to be modified. |

## New Features

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |
| 1 |  |  |  |

# Version 1.00

The release date of this version is February 2, 2011.

## Defects

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |
| 1 |  | N | Modified the firmware to apply the reversed coil correction to the velocity sum parameter rather than the engineering units in fps. This will allow the test system to use the same limits for both a sensor built correctly and one with a reversed coil. |

## Improvements

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |

## New Features

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |
| 1 |  |  |  |

# Version 0.13

The release date of this version is January 23, 2011.

## Defects

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |
| 1 |  |  |  |

## Improvements

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |

## New Features

| Item # | Track | Ext. (Y/N) | Description |
| --- | --- | --- | --- |
| 1 |  | N | Added a buffering mechanism to facilitate the collection of individual velocity sum measurements by the FH950 Sensor Test App. This feature is intended for FSBU engineering to develop filtering algorithms for velocity. |
| 2 |  | N | Implemented the calibration correction polynomials developed by Chris Tometich that are applied after the conversion from ADC to EU. It is only applied on 2pt velocity tables. |
| 3 |  | N | Updated the boot version number from v0.02 to v1.00 in preparation for product launch. The two boot versions are essentially the same except for the version number. |