PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

Single or multi-stream online Alkalinity Analyser

Analyser for automatic measurement of Alkalinity in water. The method is acid-base titration with hydrochloric acid. Reagents are dosed by high precision micropumps. The analyser performs automatic cleaning and validation. Up to 8 process streams can be monitored. Single sample ("grab sample") measurement is possible.

Technical Data

| Measurement method | Acid-base titration with hydrochloric acid |
| --- | --- |
| Measuring range | See Scope of delivery section |
| Precision | Better than 2% full scale range for standard test solutions |
| Detection limit | ≤ 10 mg/L |
| Cycle time | 10 - 15 minutes |
| Automatic cleaning | Yes |
| Calibration | N.A. |
| Validation | Automatic; frequency freely programmable |
| Ambient temperature | 10 - 30 °C ± 4 °C deviation at 5 - 95% relative humidity (non-condensing) |
| Reagent requirements | Keep between 10 - 30 °C |
| Sample pressure | By external overflow vessel |
| Flow rate | 100 - 300 mL/min |
| Sample temperature | 10 - 30 °C |
| Sample quality | Maximum particle size 100 µm, ≤ 0.1 g/L; Turbidity ≤ 50 NTU |
| Power | 100 - 240 VAC, 50/60 Hz Max. power consumption: 120 VA |
| Instrument air | Dry and oil free according to ISA-S7.0.01-1996 quality standard for instrument air |
| Demineralised water | For rinsing / dilution |
| Drain | Atmospheric pressure, vented, min. Ø 64 mm |
| Earth connection | Dry and clean earth pole with low impedance (≤ 1 Ohm) using an earth cable of Ø 2.5 mm² |
| Analogue outputs | Active 4 - 20 mA max. 500 Ohm load, standard 1, max. 8 (option) |
| Digital outputs | Optional: RS232, Modbus (TCP/IP, RS485) |
| Alarm | 1 x malfunctioning, 4 x user-configurable, max. 24 VDC/0.5 A, potential free contacts |
| Protection class | Analyser cabinet: IP55 / Panel PC: IP65 |
| Material | Hinged part: Thermoform ABS, door: plexiglass Wall section: Galvanised steel, powder coated |
| Dimensions ( H x W x D) | 690 mm x 465 mm x 330 mm |
| Weight | 25 kg |
| Certifications | CE compliant / UL certified |

Scope of delivery

[ ] (delete what is not needed)

Analytical Instrument

Online analyser for

[ ] Alkalinity, total

[ ] Alkalinity, free

Measuring Range

[ ] 10 - 500 mg/L CaCO3

[ ] 25 - 1250 mg/L

[ ] 50 - 2500 mg/L

[ ] 100 - 5000 mg/L

[ ] internal dispenser dilution (max. factor 100)

Process Streams/Channels

[_____] streams (fill in, select up to 8)

Accessories and options

[_____]x 4-20 mA outputs (fill in, select up to 8)

[_____]x 4-20 mA Outputs and Modbus RS485 (fill in, select up to 4)

[_____]x 4-20 mA Outputs and Modbus TC/IP (fill in, select up to 4)

[ ] Modbus RS485

[ ] Modbus TCP/IP

[ ] Filtration System, for immersion, pore size [_____] µm (fill in, select from 50/100/200 µm)

[ ] Filtration System, fast loop, pore size [_____] µm (fill in, select from 50/100/200 µm)

[ ] Table stand

[ ] Floor stand

Services

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:	Hach

Product: 	EZ4000 Alkalinity Analyser

DOC353.52.35246.Nov20