PLEASE NOTE: The following specification contains areas, highlighted in yellow and with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

- GENERAL

  - Section includes:

    - Instrument for semi-continuous, online monitoring of Chloride in water using titration.

  - Measurement Procedures

- The analytical method to determine Chloride in water is potentiometric titration with silver nitrate, conform with standard method ASTM 4500-Cl (B).

  - Alternates

    - Analyzers without standard automatic procedures for validation and cleaning are not acceptable.

    - Analyzers without the ability for single sample ("grab sample") measurement are not acceptable.

    - Analyzers without option for up to 8 sample streams are not acceptable.

    - Analyzers without analog and digital output options are not acceptable.

    - Analyzers without an integrated industrial panel PC are not acceptable.

  - System Description

    - Performance Requirements

      - Measuring Range

- [   ] 5 to 50 mg/L (parts per million) 
[   ] 5 to 125 mg/L  
[   ] 5 to 250 mg/L 
[   ] 25 to 500 mg/L

      - Limit of detection

        - ≤ 5 mg/L

      - Precision

        - Better than 2% full scale range for standard test solutions

    - Other Specifications

      - Cycle Time

        - 10 – 15 minutes

      - Cleaning

        - Automatic; frequency freely programmable

      - Calibration

        - NA

      - Validation

        - Automatic; frequency freely programmable

      - Alarm

        - 1x malfunctioning, 4x user-configurable, max. 24 VDC/0.5 A, potential free contacts

      - Protection class

        - Analyzer cabinet: IP55

        - Panel PC: IP65

      - Material

        - Hinged part: Thermoform ABS, door: plexiglass

        - Wall section: Galvanized steel, powder coated

      - Power Supply

        - 100 – 240 VAC, 4 A, 50/60 Hz

        - Max. power consumption: 150 VA

  - Certifications

    - CE compliant

    - UL certified

  - Environmental Requirements

    - Operational Criteria

      - Operating temperature: 10 to 30 °C ±4 °C deviation (50 to 86 °F ±7.2 °F deviation)

      - Relative humidity: 5 to 95 %, non-condensing

      - Reagent temperature: keep between 10 to 30 °C (50 to 86 °F)

  - Warranty

    - Warranted from manufacturer defects for two years (Europe) or one year (all other geographies) from date of shipment.

  - Maintenance and Service

    - Unscheduled Maintenance

      - Check and clean the analyzer components, depending on cleanliness of the sample

    - Scheduled Maintenance / preventative

      - Monthly

        - Reagents refill; validation and / or calibration

      - Quarterly

        - Pump tubing replacement

      - Annually

        - Calibration

        - Replacement of all tubing

        - Replacement of valves and pistons

PRODUCTS

  - Manufacturer

    - Hach

    - EZ4000 Series Chloride Analyzer

  - Manufactured Unit

    - The Chloride analyzer consists of a microprocessor controlled volumetric analyzer designed to monitor Chloride semi-continuously in a sample stream. Automatic cleaning and validation are available.

  - Equipment

    - Online Analyzer

      - Utilizes volumetric measurement with a high-precision micropumps for reagent dosing.

      - With automatic validation, priming and cleaning

      - If chosen, the instrument can monitor up to 8 sample streams.

      - Plug-and-play airtight fittings for connecting reagent tubing

      - High precision micropumps for reagent dosing

    - Controller

      - Industrial panel PC with 5.7” TFT color display, compact flash memory, modular compact I/O system

      - User interface with different user levels (Automatic, User Level 1, User Level 2, Administrator)

      - If chosen, includes capability to communicate measurements via 4-20 mA outputs, Modbus TCP/IP, Modbus RS485.

    - Reagents and Standards

      - The analyzer shall use quick connect reagent containers with pre-installed tubing.

      - Reagents and standards shall be prepared according instructions on the Method + Reagent sheet.

  - Components

    - Analytical instrument
To deliver:

      - Chloride Analyzer as selected in section 1.1.A.

      - Wall-mount bracket

      - Reagent containers

      - User Manual

      - Method + Reagent Sheet

    - Dimensions: refer to analyzer drawings

    - Weight: 25 kg (55 lb)

  - Instrument Options,

- *******************************************************

- Must be added to instrument at time of order.

- *******************************************************

- Number of sample streams (fill in, select up to 8):

- [     ] sample streams

- Outputs (select / fill in one)

- [     ]x 4-20 mA Outputs (fill in, select up to 8)

- [     ]x 4-20 mA Outputs and Modbus RS485 (fill in, select up to 4)

- [     ]x 4-20 mA Outputs and Modbus TC/IP (fill in, select up to 4)

- [   ] Modbus RS485

- [   ] Modbus TCP/IP

  - Instrument Accessories

- ***************************

- Select as many as required

- ***************************

- [   ] External dilution unit

- [   ] Universal Filtration System, pore size 50 µm

- [   ] Universal Filtration System, pore size 100 µm

- [   ] Heavy-Duty Filtration System, pore size 50 µm

- [   ] Heavy-Duty Filtration System, pore size 100 µm

- [   ] Table Stand

- [   ] Floor Stand

EXECUTION

  - Preparation

      - Mounting

        - As shown on the drawings

      - Inlet and outlet connection sizes

        - As shown on the drawings

      - Sample Flow Rate

        - 100 to 300 mL/minute

      - Sample Pressure

        - By external overflow vessel

      - Sample Temperature

        - 10 to 30 °C ±4 °C deviation (50 to 86 °F ±7.2 °F deviation)

      - Other sample requirements

        - Maximum particle size 100 µm, <0.1 g/L suspended solids, turbidity <50 NTU

      - Instrument air

        - Dry and oil free according to ISA-S7.0.01-1996 quality standard for instrument air

      - Demineralized water

        - For rinsing and/or dilution

      - Drain

        - Atmospheric pressure, vented, min. ø 64 mm

      - Earth connection

        - Dry and clean earth pole with low impedance (> 1 Ohm) using an earth cable of > 2.5 mm2

      - Power supply

  - Installation

    - Install analyzer following transmittal drawings and instrument user manual.

  - Manufacturer’s Service and Start-Up

    - Contractor will include the manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

    - Contractor will include a manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment startup through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

    - Items A and B are to be performed by manufacturer’s factory-trained service personnel. Field service and factory repair by personnel not employed by the manufacturer is not allowed.

    - Use of manufacturer’s service parts is required. Third-party parts are not approved for use.

DOC353.53.30533.Jan21

END OF SECTION