PLEASE NOTE: The following specification contains areas, highlighted in yellow and with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

GENERAL

  - Section includes:

    - Instrument for continuous, online monitoring of Total Organic Carbon, Total Carbon, Total Inorganic Carbon, Biological Oxygen Demand and/ or Chemical Oxygen Demand via correlation and Volatile Organic Carbon via calculation in water.

    - The analyzer has internal sensors and utilizes self-diagnostics to monitor analyzer status and produces alerts when maintenance or service are required.

  - Measurement Procedures

- The method of measuring Total Organic Carbon will be by Two Stage Advanced Oxidation which utilizes the Hydroxyl radical and a manganese catalyst.

  - Alternates

    - Other instruments that do not use the Two Stage Advanced Oxidation which utilizes the Hydroxyl radical and a manganese catalyst are not acceptable

    - Other instruments that do not have predictive diagnostic capabilities and don’t generate alerts when maintenance or service are required are unacceptable.

  - System Description

    - Performance Requirements

          - Measurement range:

- 0 – 5mgC/l,

          - Detection Limit: 0.010mgC/l

          - Repeatability:

- ±2% of reading or ±0.01mgC/l whichever is greater in 0 – 5mgC/l range,

          - Cycle time: from 5 min, depending on range and application

          - Sample Flow: minimum 100ml per analysis

          - Sample Pressure: Ambient

          - Sample Temperature: 2 to 60 °C (36 to 140 °F)

  - Certifications

    - CE Compliant

    - KC (Korea Certification)

    - ACMA (Australia & New Zealand EMC Compliance)

    - China RoHS

  - Environmental Requirements

    - Operational Criteria

      - Operating Temperature: 0 to 45 °C (32 to 113 °F)

      - Relative Humidity: 5 to 85 %, non-condensing

  - Maintenance Service

    - Unscheduled Maintenance

      - Replace chemicals as required

  - Warranty

    - Warranted for one year from date of shipment against manufacturer defects.

PRODUCTS

  - Manufacturer

    - Hach Ireland Ltd

    - Model Hach BioTector B3500ul Online TOC Analyzer

  - Analyzer

    - The online TOC analyzer consists of:

      - Housing: Glass Reinforced Polyester

      - Sampling system

      - Oxygen concentrator

      - Ozone generator

      - Oxidation reactor

      - Non-dispersive Infrared (NDIR) analyzer

  - Equipment

    - Online TOC analyzer

      - To measure TOC of the sample it utilizes hydroxyl radicals as the oxidizing agent.

      - The TOC analyzer automatically compensates for temperature and pressure utilizing an embedded temperature and pressure sensors.

      - Includes capability to actively monitor internal components and present diagnostics on the overall health of the TOC analyzer and time to next required preventive maintenance.

      - Includes capability to actively monitor reagents usage, displays current reagent volumes and present approximate time to the next reagents replacement.

      - Data transmission is made with a SD card.

  - Components

    - Standard equipment:

      - Analyzer

      - Mounting brackets

      - User Manual

      - Reagent dip tubes

    - Dimensions: Refer to TOC analyzer drawings

    - Weight: typically, 110 lbs (50 kg)

  - Instrument Options,

- *******************************************************

- Must be selected at time of order. Choose one

- *******************************************************

[ ] Single channel Online Total Organic Carbon Analyzer, 0 - 5 mg/L C,

[ ] Dual channel Online Total Organic Carbon Analyzer, 0 - 5 mg/L C

[ ] Single channel Online Total Organic Carbon Analyzer, IP54 Purge Ready, 0 - 5 mg/L C,

[ ] Dual channel Online Total Organic Carbon Analyzer, IP54 Purge Ready, 0 - 5 mg/L C

[ ] Single channel Online Total Organic Carbon Analyzer, ATEX Compliant, EExp Zone 2, 0 - 5 mg/L C,

[ ] Dual channel Online Total Organic Carbon Analyzer, ATEX Compliant, EExp Zone 2, 0 - 5 mg/L C,

[ ] Single channel Online Total Organic Carbon Analyzer, ETL Certified, EExp Class 1 Div 2, 0 - 5 mg/L C,

[ ] Dual channel Online Total Organic Carbon Analyzer, ETL Certified, EExp Class 1 Div 2, 0 - 5 mg/L C

  - Instrument Accessories

- ********************

- Select as many as required

- ********************

- [ ] VOC mode used for measurement of  Volatile Organic Carbon in the sample

- [ ] Grab sample/Calibration port

- [ ] Remote activation of Manual/Calibration port

- [  ] Sample Sensor to detect presence of the sample and evaluate the quantity of air bubbles in the sample

- [  ] BioTector compressor

- [  ] Filter pack for removal of water, dirt, and oil from compressed air (if using instrument air supply)

Choose one or none

- [  ] Profibus

- [  ] Modbus RTU

- [  ] Modbus TCP/IP

Choose one or none

- [  ] BioTector Sample Overflow Chamber

- [  ] BioTector Flow Through Sampling Chamber with Sand Trap

- [  ] BioTector Sampling Chamber with Sand Trap

EXECUTION

  - Preparation

      - Mounting

        - As shown on the drawings

      - Inlet and outlet connection sizes

        - As shown on the drawings

  - Installation

    - Install TOC analyzer following transmittal drawings and instrument user manual.

  - Manufacturer’s Service and Start-Up

    - Contractor will include the manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

    - Contractor will include a manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment startup through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

    - Items A and B are to be performed by manufacturer’s factory-trained service personnel.  Field service and factory repair by personnel not employed by the manufacturer is not allowed.

    - Use of manufacturer’s service parts and reagents is required.  Third-party parts and reagents are not approved for use.

- END OF SECTION