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1.1 Section includes

A. High Output Air Blast (HOAB). Compressed air will be provided to make possible to clean a probe . A probe specific cleaning head needs to be used to fit with the specific probe involved. Probe specific cleaning head is not a part of the HOAB.

1.2 Measurement Procedures

A. Compressed air generated from a compressor will be let to a cleaning device (cleaning head). The design of the probe specific cleaning head makes sure that the surface of the measuring part of the probe remains clean.
 
1.3 Alternates

A. Manual cleaning of the measuring part of the probe are not acceptable

1.4 System Description

A. Performance Requirements
1. Air pressure (at compressor outlet): 115VAC 3.1 bar (45psi) / 230VAC 2.8 bar (40 psi) 
2. Air flow (at compressor outlet): 115VAC 2.1 m3/hour / 230VAC 1.8m3/hour

1.5 Certifications

A. UL, CSA, cETLus, CE

1.6 Environmental Requirements

A. Operational Criteria
1. Operating temperature -20 to 50 °C (-4 to 122 °F) 95% rel. humidity non-considering

1.7 Warranty

A. The product includes a one-year warranty from date of shipment. (excluding wear parts)

1.8 Maintenance Service

A. Scheduled maintenance: 
1. Weekly / Monthly: visual inspection, if necessary, cleaning of  the air intake, depending on location




PRODUCTS

1.9 Manufacturer

A. Hach Company, Loveland, CO
1. HOAB, High Output Air Blast system

1.10 Manufactured Unit

A. The HOAB -consist out of a non-metallic cabinet (enclosure) including a compressor- which needs be installed along the probe equipped with a cleaning head at the basin and approx. 9m (27ft) air hose.

1.11 Equipment
A. The HOAB is a stand-alone system (AC powered),
including a controller relay fuse (T, 0.25 A, 250 V).
B. The enclosure is made of Polycarbonate and is suitable for outdoor installation. 
C. The air hose is suitable for outdoor installation.

1.12 Components

A. Standard equipment: 
1. Control unit
Polycarbonate
2. Air hose
 Tygon style, braided
3. Mounting accessories
 Stainless Steel 
4. Manual

B. Dimensions:
1. Control unit: (W x H x D) 351.6 x 186.9 x 409.7 mm (13.84 x 7.36 x 16.13 inches)
C. Weight:
1. Control unit: 11.2 kg 24.7 lbs 

1.13 Accessories

A. Probe specific cleaning head (kits) 












EXECUTION

1.14 Preparation

The HOAB enclosure module must be mounted steady on side of the basin. The HOAB must be installed nearby the probe with the cleaning head.

1.15 Installation

A. Contractor will install the probe in strict accordance with the manufacturer’s instructions and recommendation.
B. Manufacturer’s representative will include a half-day of start-up service by a factory-trained technician, if requested.
1. Contractor will schedule a date and time for start-up.
2. Contractor will require the following people to be present during the start-up procedure. 
a. General contractor
b. Electrical contractor
c. Hach Company factory trained representative
d. Owner’s personnel
e. Engineer
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