GENERAL

  - Section includes

    - High Output Air Blast (HOAB). Compressed air will be provided to make possible to clean a probe . A probe specific cleaning head needs to be used to fit with the specific probe involved. Probe specific cleaning head is not a part of the HOAB.

  - Measurement Procedures

    - Compressed air generated from a compressor will be let to a cleaning device (cleaning head). The design of the probe specific cleaning head makes sure that the surface of the measuring part of the probe remains clean.

  - Alternates

    - Manual cleaning of the measuring part of the probe are not acceptable

  - System Description

    - Performance Requirements

      - Air pressure (at compressor outlet): 115VAC 3.1 bar (45psi) / 230VAC 2.8 bar (40 psi)

      - Air flow (at compressor outlet): 115VAC 2.1 m3/hour / 230VAC 1.8m3/hour

  - Certifications

    - UL, CSA, cETLus, CE

  - Environmental Requirements

    - Operational Criteria

      - Operating temperature -20 to 50 °C (-4 to 122 °F) 95% rel. humidity non-considering

  - Warranty

    - The product includes a one-year warranty from date of shipment. (excluding wear parts)

  - Maintenance Service

    - Scheduled maintenance:

      - Weekly / Monthly: visual inspection, if necessary, cleaning of  the air intake, depending on location

PRODUCTS

  - Manufacturer

    - Hach Company, Loveland, CO

      - HOAB, High Output Air Blast system

  - Manufactured Unit

    - The HOAB -consist out of a non-metallic cabinet (enclosure) including a compressor- which needs be installed along the probe equipped with a cleaning head at the basin and approx. 9m (27ft) air hose.

  - Equipment

    - The HOAB is a stand-alone system (AC powered),
including a controller relay fuse (T, 0.25 A, 250 V).

    - The enclosure is made of Polycarbonate and is suitable for outdoor installation.

    - The air hose is suitable for outdoor installation.

  - Components

    - Standard equipment:

      - Control unit

- Polycarbonate

      - Air hose

- Tygon style, braided

      - Mounting accessories

- Stainless Steel

      - Manual

    - Dimensions:

      - Control unit: (W x H x D) 351.6 x 186.9 x 409.7 mm (13.84 x 7.36 x 16.13 inches)

    - Weight:

      - Control unit: 11.2 kg 24.7 lbs

  - Accessories

    - Probe specific cleaning head (kits)

EXECUTION

  - Preparation

- The HOAB enclosure module must be mounted steady on side of the basin. The HOAB must be installed nearby the probe with the cleaning head.

  - Installation

    - Contractor will install the probe in strict accordance with the manufacturer’s instructions and recommendation.

    - Manufacturer’s representative will include a half-day of start-up service by a factory-trained technician, if requested.

      - Contractor will schedule a date and time for start-up.

      - Contractor will require the following people to be present during the start-up procedure.

        - General contractor

        - Electrical contractor

        - Hach Company factory trained representative

        - Owner’s personnel

        - Engineer

- END OF SECTION