PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

# Online Ammonium Analyser with Predictive Diagnostics

Analyser for monitoring of NH4-N and NH4 in water. The method of measuring is by gas sensing electrode (GSE) that uses liquid to gas-phase conversion. The analyser calibrates and cleans itself automatically. Is capable of measuring as frequently as every 5 minutes and is equipped with self-diagnostic routines. Includes a humidity sensor to detect leakage and automatically initiate a safe shutdown. The analyser requires a universal digital controller.

The analyser has built in predictive diagnostics capability which increases available operation time. The predictive diagnostics produces alerts for upcoming maintenance tasks by monitoring the instrument’s internal components and tracking service requirements. When connected to a properly equipped controller the controller will display status indicators for each sensor with predictive diagnostics capability.

Includes the capability to remotely monitor sensors on any browser-enabled device and present diagnostics on the overall health of the measurements (on predictive diagnostics-enabled sensors), as well as upcoming and required maintenance - reducing user risk and downtime.

# Technical Data

| Measuring range | (1) 0.02 - 5.0 mg/L NH4-N |
| --- | --- |
|  | (2) 0.05 - 20 mg/L NH4-N |
|  | (3) 1 - 100 mg/L NH4-N |
|  | (4) 10 - 1000 mg/L NH4-N |
| Detection limit | Measuring range 1: 0.02 mg/L NH4-N Measuring range 2: 0.05 mg/L NH4-N Measuring range 3: 1 mg/L NH4-N Measuring range 4: 10 mg/L NH4-N |
| Accuracy | With Standard Solution: Measuring range 1: ≤ 1 mg/L: 3% of the measured value + 0.02 mg/L; > 1 mg/L: 5% of the measured value + 0.02 mg/L Measuring range 2: 3% of the measured value + 0.05 mg/L Measuring range 3: 3% of the measured value + 1.0 mg/L Measuring range 4: 4.5% of the measured value + 10 mg/L |
| Repeatability | With Standard Solution: Measuring range 1: 3% of the measured value + 0.02 mg/L Measuring range 2: 2% of the measured value + 0.05 mg/L Measuring range 3: 2% of the measured value + 1.0 mg/L Measuring range 4: 2% of the measured value + 10 mg/L |
| Response time | Measuring range 1: ≤ 0.2 mg/L: 3 measurements required (15 total minutes minimum); > 0.2 mg/L: 1 measurement required (5 total minutes) Measuring range 2: < 5 minutes Measuring range 3: < 5 minutes Measuring range 4: < 5 minutes |
| Measurement method | Gas sensing electrode (GSE) |
| Measuring interval | 5, 10, 15, 20 or 30 min. (user adjustable) |
| pH range | 5 - 9 pH |
| Permissible Chloride range | ≤ 1000 mg/L Cl- |
| Sample quality | Ultra filtrated or comparable (integrated sample filtration < 0.45 µm) |
| Flow rate | 1.0 - 20.0 L/h to the analyser from the external sample filtration |

| Sample pressure | 2.5 bar (0.25 mPa) maximum |
| --- | --- |
| Sample temperature | 4 - 40 °C |
| Special remarks | The liquid level in the basin must be below the bottom of the analyser |
| Operating temperature range | -20 - 45 °C; 95% relative humidity, non-condensing |
| Storage conditions | -20 - 60 °C; 95% relative humidity, non-condensing |
| Power supply | Power supplied by digital controller (depending on controller type an additional power box may be needed). Analyser and heated drain tubing (powered by analyser): 115 VAC or 230 VAC |
| Cable length | Data and power cable: 2 m from edge of enclosure |
| Material | PUR 66 UV-resistant |
| Enclosure waterproof rating | Rating: IP55, NEMA 3R UL50E Material: PUR 66 |
| Controller compatibility | Digital Controller |
| Dimensions | 575 x 991 x 425 mm (22.63 x 39.01 x 16.73 in.) |
| Weight | Approximately 45 kg without chemicals |

Scope of delivery

[ ] (delete what is not needed)

# Analytical Instrument

[ ] Online Ammonium Analyser for use with external filtration system

[ ] Online Ammonium Analyser with sample supply pump for use with integrated filtration system

[ ] Online Ammonium Analyser with sample supply pump for use with integrated filtration system, rough samples

# Number of Channels

[ ] 1 channel [ ] 2 channels

# Instrument Options

[ ] Flow detection

[ ] Flow detection + automatic grab sample and sample verification

# Accessories

[ ] Keys for analyser enclosure (1 replacement pair)

[ ] Heated drain/connecting hose, 2 m, 230 V

[ ] Mounting kit for analyser incl. fastening and angle bracket, screws [ ] Accessories for continuous sample (1-/2- channel)

[ ] Power extension cable for digital controller, 5 m, 115-230 V AC [ ] Door for analyser indoor housing including. 4 stickers

[ ] Door for analyser enclosure including 4 instrument labels [ ] Power supply box for analyser, with EU plug

[ ] Power connection box for analysers

# Mounting Hardware

[ ] Extension pipe, 1.0 m, made of stainless steel [ ] Extension pipe, 1.8 m, made of stainless steel [ ] Rail mounting for analyser with controller

[ ] Stand mounting kit, suitable for 1 analyser + 1 digital controller [ ] Stand mounting kit, suitable for 1 analyser

[ ] Rail mounting for analyser without digital controller

[ ] Weather guard for digital controller for outdoor installation

# Services

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

[ ] External sample filtration system for Ammonium Analyser

NOTE: Sample filtration is strongly recommended in all waste water process areas. If you are specifying a two stream analyser in this design you should specify two filtration units, one for each sample stream!

Brand:	Hach

Product:	NH6000sc Online Ammonium Analyser with Prognosys Predictive Diagnostics